MODULAR SMART INFANT KIT

Retrofit Architecture — Engineering & Product Specification

Document Ref: MSIC-SPEC-2026-V2.1

Status: Finalized for Development

Distribution: Engineering, Firmware, Embedded AI, and Product Teams

Inventors: Vejay Sundaram Rajkumar  |  Nethra Nandakumar  |  Skanda A Sunder

# Table of Contents

Table of Contents	1

1. Strategic Mission & System Matrix	1

2. Volume Manufacturing Cost Architecture	1

2.1 Consumer Market Reference BOM	1

2.2 Clinical Mission-Critical Reference BOM	1

3. Physical Hardware Pinout & Bus Logic	1

3.1 Pin Assignment Table	1

3.2 Electrical Driver Implementation Criteria	1

4. Clinical-Grade False-Alarm Filters	1

4.1 Filter A — Hygiene Differential Baseline Matrix	1

4.2 Filter B — Multi-Modal Distress Event Fusion	1

5. Open Items for Engineering & Product Sign-off	1

# 1. Strategic Mission & System Matrix

To satisfy strict clinical/institutional safety criteria while preserving optimized consumer hardware margins, this ecosystem splits into two drop-in hardware reference configurations that share an identical underlying software framework and algorithmic core.

| Logical Vector | Consumer Market Spec | Clinical NICU Spec | Owning Dev Team |
| --- | --- | --- | --- |
| Edge AI Vision Subsystem | Panasonic AMG8833 Grid-EYE 8×8 thermopile array (64 zones) 60° symmetric lens FOV Target: occupancy & restlessness detection | Melexis MLX90640-D55 32×24 thermopile array (768 zones) 55° × 35° narrow lens FOV Target: anatomical core segmentation | Vision / ML Core Team |
| Edge Compute Node | ESP32-S3 WROOM-1 MCU Dual-core Xtensa LX7 8MB Flash / 2MB PSRAM | Raspberry Pi Zero 2 W SoC Quad-core 64-bit ARM 512MB LPDDR2 SDRAM | Firmware & OS Team |
| Edge ML Runtime | TensorFlow Lite Micro ESP-NN vector acceleration | ONNX Runtime Lite Engine ARM NEON SIMD hardware layer | Embedded AI Team |
| Kinematics Subsystem | InvenSense MPU6050 6-axis digital accel/gyro IMU Target: structural cradle tilt alerts | PVDF piezoelectric film matrix + MPU6050 Under-mattress ribbon array insert Target: contactless pulse & respiration | Signal Processing Team |
| Hygiene Telemetry | ScioSense ENS160 gas module Digital TVOC / NH3 cross-equivalents Target: local diaper event ingress | Dual ScioSense ENS160 array Multi-gas VOC tracking Target: differential clinical air check | Hardware Assembly Team |
| Environmental Context | AHT20 integrated I²C chip Accuracy: ±0.3°C / ±2% RH Target: differential room baseline | AHT20 integrated I²C chip Accuracy: ±0.3°C / ±2% RH Target: incubator enclosure validation | Firmware Core Team |
| Acoustic Audio Core | INMP441 I2S digital MEMS mic SNR: 61 dBA / flat 60Hz–15kHz Target: cry pattern capture & matching | INMP441 I2S digital MEMS mic SNR: 61 dBA / flat 60Hz–15kHz Target: high-fidelity acoustic diagnostic logs | Audio Pipeline Team |

# 2. Volume Manufacturing Cost Architecture

Production sourcing calculations are based on high-volume contract manufacturing scaling limits (1k–10k component lot runs).

## 2.1 Consumer Market Reference BOM

| Component | Unit Cost (₹, volume pricing) |
| --- | --- |
| Panasonic AMG8833 thermal array | ₹1,600 – ₹1,900 |
| ESP32-S3 WROOM-1 microcontroller | ₹250 – ₹300 |
| InvenSense MPU6050 6-axis IMU | ₹70 – ₹110 |
| ScioSense ENS160 digital gas sensor | ₹180 – ₹240 |
| AHT20 temperature/humidity module | ₹45 – ₹65 |
| INMP441 I2S digital microphone node | ₹65 – ₹95 |
| TOTAL — Consumer raw hardware BOM | ~₹2,210 – ₹2,710 |

# 3. Physical Hardware Pinout & Bus Logic

For immediate trace-routing by electrical schematic design teams (consumer/ESP32-S3 reference build).

┌────────────────────────────────────┐ │ ESP32-S3 WROOM-1 MCU │ │ │ INMP441 Mic ────>│ GPIO12 (I2S_WS) │ │ GPIO13 (I2S_SD) │ │ GPIO14 (I2S_SCK) │ │ │ AMG8833 Thermal >│ GPIO41 (SDA) ── Shared I2C Bus │ ENS160 Gas >│ GPIO42 (SCL) ── (400 kHz, pull-ups) │ AHT20 Env >│ │ │ │ MPU6050 IMU >│ GPIO4 (IMU Interrupt Event Pin) │ └────────────────────────────────────┘

## 3.1 Pin Assignment Table

| Peripheral | ESP32-S3 Pin | Bus / Function |
| --- | --- | --- |
| INMP441 I2S microphone | GPIO12 | I2S_WS (Word Select) |
| INMP441 I2S microphone | GPIO13 | I2S_SD (Serial Data) |
| INMP441 I2S microphone | GPIO14 | I2S_SCK (Serial Clock) |
| AMG8833 thermal array | GPIO41 | I²C SDA — shared bus, 400 kHz, 4.7kΩ pull-up |
| ENS160 diaper sensor | GPIO41 / GPIO42 | I²C SDA/SCL — shared bus |
| AHT20 room sensor | GPIO41 / GPIO42 | I²C SDA/SCL — shared bus |
| MPU6050 6-axis IMU | GPIO42 | I²C SCL — shared bus |
| MPU6050 interrupt line | GPIO4 | Edge-triggered hardware interrupt (bypasses polling) |

## 3.2 Electrical Driver Implementation Criteria

- I²C shared telemetry bus: GPIO41 (SDA) and GPIO42 (SCL) on the ESP32-S3. Requires dedicated 4.7kΩ pull-up resistors on the board layout, running at standard Fast-Mode frequency (400 kHz).

- Fixed hex address registers — see table below.

- I²S digital audio bus routing: microphone clock and line-out connect directly to GPIO12 (WS), GPIO13 (SD), and GPIO14 (SCK).

- Hardware interrupt lines: MPU6050 INT pin connects directly to GPIO4. Firmware must configure this as an edge-triggered input interrupt to catch structural mechanical displacement instantly, without standard polling-queue delays.

### Fixed I²C Address Registers

| Device | Fixed I²C Hex Address |
| --- | --- |
| Panasonic AMG8833 (thermal) | 0x69 |
| ScioSense ENS160 (gas) | 0x53 |
| AHT20 (environment) | 0x38 |
| InvenSense MPU6050 (IMU) | 0x68 |

# 4. Clinical-Grade False-Alarm Filters

To prevent institutional alarm fatigue and consumer attrition, developers must strictly enforce the following validation filters.

## 4.1 Filter A — Hygiene Differential Baseline Matrix

The gas sensor processes diaper events by tracing gas-density shifts relative to room atmospheric fluctuations, overriding legacy raw absolute-threshold triggers.

def verify_diaper_event(mattress_tvoc_ppb, ambient_tvoc_ppb): # Evaluate gas density deviation against the nursery room baseline delta_hygiene_threshold = mattress_tvoc_ppb - ambient_tvoc_ppb if delta_hygiene_threshold > CRITICAL_DIAPER_SPIKE_LIMIT: return True # Validated diaper event trigger approved return False # Spurious noise anomaly flagged; event rejected

## 4.2 Filter B — Multi-Modal Distress Event Fusion

Critical alerts must bypass normal background runtime threads. High-priority alerts require concurrent confirmation across distinct sensory nodes before a distress packet is dispatched.

[ TFLite Audio CNN ] ──> Crying Probability > 90% ──┐ ├─> [ SYSTEMIC DISTRESS ALERT ] [ ONNX Spatial CNN ] ──> Posture == Restless ──┤ • Bypass background queues │ • Immediate edge-to-app push [ Mechanical PVDF ] ──> Rhythmic Pulse == None ──┘

This multi-tier verification protects clinical NICU neonates against system dropouts while keeping the consumer model robust and cost-efficient at scale.

# 5. Open Items for Engineering & Product Sign-off

- Confirm whether the consumer and clinical variants ship on a single firmware image with a config flag, or as separate builds — this affects the CI/CD branching strategy.

- Define CRITICAL_DIAPER_SPIKE_LIMIT and per-sensor calibration constants as part of firmware config, not hardcoded values.

- Clinical variant (PVDF pulse/respiration + Raspberry Pi Zero 2 W) will likely fall under stricter regulatory review (e.g., medical device / IEC 62304-adjacent) — Product to confirm target certification path before hardware freeze.

- Confirm data retention and privacy handling for acoustic and thermal data given India's DPDP Act 2023, especially for the clinical/hospital deployment model.

- Reconcile domain naming ('AC for Common Man' vs. 'Smart Cradles') flagged in the wider portfolio review before this spec is finalized for the CFO/advisory panel.